1. Device Description

A pulse oximeter is a non-invasive, active medical device intended to measure functional arterial oxygen saturation (SpO₂) and pulse rate by analyzing the absorption characteristics of red and infrared light as they pass through pulsatile vascular tissue. The device provides real-time physiological monitoring and is widely used to assess a patient's oxygenation status and circulatory function in both clinical and non-clinical healthcare environments.

The operating principle of pulse oximetry is based on spectrophotometric measurement and photoplethysmography. Oxygenated hemoglobin and deoxygenated hemoglobin absorb light differently at specific wavelengths. The pulse oximeter sensor typically contains two light-emitting diodes (LEDs), one emitting red light (approximately 660 nm) and the other emitting infrared light (approximately 940 nm), along with a photodetector. During operation, the emitted light passes through a vascular tissue bed, such as a fingertip, toe, or earlobe. The photodetector measures the amount of transmitted light reaching the sensor, and proprietary algorithms analyze the pulsatile component of blood flow to calculate arterial oxygen saturation and pulse rate.

The device generally consists of a sensor assembly, signal acquisition circuitry, microprocessor-based processing unit, software algorithms, display interface, alarm system (where applicable), and power source. Depending on the intended application, pulse oximeters may be available in several configurations, including portable fingertip devices, handheld monitors, bedside monitoring systems, wearable devices, and integrated multi-parameter patient monitoring systems. Certain models may support continuous monitoring, while others are intended for spot-check measurements.

The sensor component may be reusable or single-use and is designed to maintain consistent contact with the monitoring site to ensure accurate signal acquisition. Device software processes raw optical signals, compensates for interference caused by motion artifacts or low perfusion conditions, and displays measured physiological parameters. Advanced devices may include signal quality indicators, perfusion index measurements, alarm functions, data storage capabilities, wireless communication interfaces, and connectivity to hospital information systems or electronic medical records.

The pulse oximeter is intended to provide rapid, continuous, and non-invasive assessment of oxygen saturation without the need for arterial blood sampling. It is commonly used in hospitals, operating rooms, intensive care units, emergency departments, outpatient clinics, ambulances, long-term care facilities, and home healthcare settings. The device supports clinical decision-making by enabling early detection of hypoxemia and facilitating the monitoring of patients receiving oxygen therapy, respiratory support, anesthesia, or treatment for cardiopulmonary conditions.

The device may be used across a broad patient population, including adults, pediatric patients, and neonates when appropriate sensors and validated measurement ranges are employed. Performance characteristics are typically established through bench testing, verification activities, and clinical validation studies comparing device measurements against recognized reference methods such as arterial blood gas analysis and co-oximetry.

As an active monitoring device, the pulse oximeter does not deliver therapy or directly influence physiological function. Instead, it provides clinically relevant information that assists healthcare professionals in evaluating patient status, monitoring disease progression, assessing treatment effectiveness, and responding to changes in oxygenation levels. The device forms an integral part of modern patient monitoring systems and is recognized as a standard tool for respiratory and cardiovascular assessment across a wide range of healthcare applications.

2. Intended Purpose

The pulse oximeter is intended to non-invasively measure and display arterial oxygen saturation (SpO₂) and pulse rate in patients. The device is designed to provide healthcare professionals and trained users with real-time physiological information that supports the assessment of a patient's respiratory and cardiovascular status. It is intended for spot-check monitoring or continuous monitoring depending on the device configuration and clinical setting.

The device may be used across a broad range of patient populations, including adults, pediatric patients, and neonates, provided appropriate sensors and accessories are utilized. Pulse oximeters are commonly employed in hospitals, outpatient clinics, emergency departments, operating rooms, intensive care units, ambulatory care settings, and home healthcare environments. The device supports clinical decision-making by assisting in the detection of hypoxemia, monitoring oxygen therapy effectiveness, and evaluating respiratory function during anesthesia, surgery, critical care, and recovery. It is intended as a monitoring aid and should be used alongside other clinical assessments and diagnostic information rather than as a sole basis for medical decisions.

3. Device Design and Development Overview

The pulse oximeter is designed as a non-invasive physiological monitoring device intended to measure arterial oxygen saturation (SpO₂) and pulse rate. The design of the device is based on the principle of spectrophotometry and photoplethysmography, which enable the detection and analysis of pulsatile blood flow through peripheral tissue. The primary design objective is to provide accurate, reliable, and real-time monitoring of oxygenation status while ensuring patient safety, ease of use, and compatibility with various clinical and non-clinical healthcare environments.

The device architecture typically consists of a sensor module, signal acquisition circuitry, microprocessor-based signal processing unit, display interface, power management system, and embedded software. The sensor module incorporates red and infrared light-emitting diodes (LEDs) and a photodetector. The LEDs emit light at predetermined wavelengths through vascular tissue, while the photodetector measures the transmitted light intensity. The acquired optical signals are converted into electrical signals and processed by embedded algorithms to calculate oxygen saturation and pulse rate values.

Design inputs for the pulse oximeter include user needs, clinical requirements, applicable regulatory requirements, safety requirements, performance specifications, and usability considerations. User requirements focus on obtaining rapid, non-invasive, and accurate oxygen saturation measurements with minimal training. Regulatory requirements include compliance with applicable medical device regulations and standards such as ISO 13485, ISO 14971, IEC 60601 series standards, IEC 62304 for software lifecycle processes, and IEC 62366 for usability engineering.

The design process incorporates risk management activities throughout development to identify potential hazards associated with device operation. Risks related to measurement inaccuracies, sensor malfunction, electrical safety, electromagnetic interference, software failure, and user error are evaluated and controlled through design features, verification testing, and labeling. Human factors and usability considerations are integrated into the design to minimize the likelihood of use errors and ensure safe operation across the intended user population.

Verification and validation activities are conducted to confirm that design outputs meet established design inputs and intended performance requirements. Verification testing includes electrical safety testing, electromagnetic compatibility testing, software verification, environmental testing, alarm testing, and performance characterization. Validation activities include usability validation and clinical performance evaluation against recognized reference methods such as arterial blood gas analysis and co-oximetry.

The final device design provides continuous or spot-check monitoring capabilities, depending on the intended configuration, and may include additional features such as visual and audible alarms, data storage, wireless communication, battery operation, trend analysis, and integration with hospital information systems. The design supports safe and effective operation across a broad range of patient populations and healthcare settings while maintaining compliance with applicable regulatory and quality requirements.

4. Manufacturing Process Overview

The manufacturing process for the pulse oximeter is controlled through an established Quality Management System (QMS) to ensure that all finished devices consistently meet predefined specifications, safety requirements, and performance standards. Manufacturing activities are conducted in accordance with documented procedures, work instructions, process controls, and quality assurance requirements established under ISO 13485 and applicable regulatory requirements.

The manufacturing process begins with the procurement and inspection of approved raw materials, electronic components, optical components, plastic housings, sensors, printed circuit boards (PCBs), display modules, batteries, and packaging materials. All incoming materials are sourced from qualified suppliers and undergo incoming inspection to verify compliance with specified quality and performance requirements before release into production.

Printed circuit board assembly is performed through controlled manufacturing processes that include automated component placement, soldering operations, and inspection activities. Electronic components such as microprocessors, memory devices, signal conditioning circuits, LEDs, photodetectors, and power management systems are assembled onto the PCB according to approved engineering specifications. Following assembly, electrical testing is performed to verify proper functionality and identify any manufacturing defects.

Sensor manufacturing involves the integration of red and infrared LEDs with photodetector components into a sensor housing designed to maintain proper optical alignment and measurement accuracy. Sensor assemblies undergo calibration and functional testing to verify optical performance and signal quality before integration into the final device. Depending on the device design, reusable or disposable sensor configurations may be manufactured.

Following assembly of the electronic and sensor components, the device is integrated into its final enclosure. Software and firmware are installed and configured according to approved software release procedures. Functional testing is then performed to verify operation of all device features, including oxygen saturation measurement, pulse rate calculation, display functionality, alarm systems, power management, and communication interfaces where applicable.

Quality control inspections are conducted throughout the manufacturing process to ensure compliance with product specifications. These inspections may include visual inspection, dimensional verification, electrical safety testing, performance testing, software verification checks, and final acceptance testing. Any non-conforming products are identified, segregated, investigated, and managed through established quality procedures.

Final product release occurs only after all manufacturing records, inspection results, and testing activities have been reviewed and approved by authorized personnel. Finished devices are packaged in validated packaging systems designed to protect the product during storage and transportation. Product labeling, unique device identification (UDI) information, instructions for use, and regulatory markings are applied as part of the final packaging process.

Manufacturing records are maintained to ensure full traceability of materials, components, production activities, inspections, testing results, and product release decisions. These records support regulatory compliance, post-market surveillance activities, complaint investigations, corrective and preventive actions (CAPA), and ongoing quality management activities throughout the product lifecycle.

The below documents have been attached to the folder :

5. Risk Management File

6. Verification Report

7. Validation Report

8. Clinical Evaluation Report

9. Labeling

Labeling Information

The labeling of the pulse oximeter has been developed to provide users with clear, accurate, and comprehensive information necessary for the safe identification, handling, storage, and operation of the device throughout its lifecycle. The labeling is designed in accordance with applicable regulatory requirements and relevant standards governing medical device information supplied by the manufacturer. All labeling information is intended to support proper device use, facilitate traceability, and communicate important safety information to healthcare professionals and end users.

The device label clearly identifies the product by its trade name, model number, and unique device identification (UDI) information. Manufacturer details, including the legal manufacturer name and address, are provided to ensure traceability and regulatory compliance. Product-specific identifiers such as serial numbers, batch numbers, and manufacturing information are included to support product tracking, complaint investigations, field corrective actions, and post-market surveillance activities when required.

The labeling communicates the intended purpose of the pulse oximeter as a non-invasive monitoring device intended for the measurement of arterial oxygen saturation (SpO₂) and pulse rate. Information regarding the intended patient population and intended users is provided to ensure that the device is utilized within its validated operating conditions. The label further identifies whether the device is intended for spot-check monitoring, continuous monitoring, or both, depending on the specific device configuration.

Safety-related information forms an important component of the labeling. Relevant warnings, precautions, and limitations are communicated to users to reduce the likelihood of misuse and measurement errors. Users are informed that pulse oximeter readings should be interpreted in conjunction with the overall clinical condition of the patient and should not be used as the sole basis for diagnosis or treatment decisions. The labeling highlights factors known to affect measurement accuracy, including patient movement, poor peripheral circulation, improper sensor placement, excessive ambient light, nail polish, artificial fingernails, and the presence of abnormal hemoglobin conditions.

Standardized symbols are incorporated into the labeling to facilitate international understanding and compliance with recognized medical device labeling conventions. Symbols relating to manufacturer identification, serial number, batch number, temperature limitations, consult instructions for use, electrical safety classification, and regulatory compliance are displayed where appropriate. The use of standardized symbols reduces language dependency and enhances user comprehension across different geographic markets.

Storage and transportation requirements are also included within the labeling. Information regarding acceptable temperature ranges, humidity limits, handling precautions, and environmental conditions is provided to ensure that the device maintains its specified performance characteristics throughout storage and distribution. These instructions support preservation of device integrity and minimize the risk of performance degradation prior to use.

In addition to identification and safety information, the labeling provides regulatory compliance details including applicable conformity markings and certification references. These markings indicate that the device has undergone the necessary conformity assessment procedures and complies with relevant regulatory requirements. The labeling therefore serves not only as a means of device identification but also as a critical component of risk communication, traceability, regulatory compliance, and user safety throughout the product lifecycle.

Overall, the labeling has been designed to ensure that users receive the information necessary to safely identify, handle, store, and operate the pulse oximeter. The information supplied through device labeling supports effective risk management and contributes to the safe and intended use of the device within its specified clinical environment.

10. Instructions for Use (IFU)

Instructions for Use Information

The Instructions for Use (IFU) document provides detailed information intended to ensure the safe, effective, and appropriate operation of the pulse oximeter by healthcare professionals and other authorized users. The IFU serves as the primary source of user guidance and contains comprehensive information regarding device functionality, intended use, operating procedures, safety considerations, maintenance requirements, and interpretation of measurement results.

The IFU describes the pulse oximeter as a non-invasive medical device designed to measure arterial oxygen saturation (SpO₂) and pulse rate through optical sensing technology. The document explains the operating principle of pulse oximetry, including the use of red and infrared light wavelengths to estimate oxygen saturation levels based on the absorption characteristics of oxygenated and deoxygenated hemoglobin. This information provides users with a basic understanding of device functionality and supports appropriate interpretation of measurement results.

The intended purpose section specifies that the device is intended to provide non-invasive monitoring of oxygen saturation and pulse rate in patients requiring assessment of oxygenation status. The document identifies the intended patient population, intended users, and intended use environments. Users are instructed to operate the device only within the conditions and patient populations for which performance has been validated. Any known limitations regarding patient characteristics, environmental conditions, or clinical applications are clearly described.

Detailed operating instructions are provided to ensure proper use of the device. The IFU describes the procedures for preparing the device, powering the system, selecting the appropriate sensor, and positioning the sensor correctly on the patient. Instructions emphasize the importance of proper sensor placement and maintaining adequate sensor contact throughout monitoring to ensure reliable signal acquisition and measurement accuracy. Guidance is also provided regarding verification of displayed measurements and recognition of signal quality indicators where applicable.

The IFU includes comprehensive safety information intended to minimize risks associated with device use. Users are informed that pulse oximetry should be used as an adjunct to, rather than a replacement for, comprehensive patient assessment. The document explains that oxygen saturation readings may be affected by factors including patient movement, poor perfusion, excessive ambient light, venous pulsations, severe hypotension, hypothermia, cardiac arrhythmias, nail polish, artificial fingernails, and abnormal hemoglobin conditions. Users are instructed to consider these factors when interpreting displayed values and to seek alternative assessment methods when clinically indicated.

Warnings and precautions are included throughout the document to address foreseeable misuse and known device limitations. Users are advised not to rely solely on pulse oximeter measurements when making clinical decisions and are instructed to evaluate the patient’s overall clinical condition in conjunction with displayed readings. The IFU further emphasizes the importance of using only manufacturer-approved sensors, accessories, and replacement components to ensure continued device performance and regulatory compliance.

Maintenance and cleaning instructions are provided to support ongoing device reliability and infection control practices. The IFU describes procedures for routine inspection of the device, cleaning of external surfaces, disinfection of reusable sensors, and replacement of damaged components. Users are instructed to inspect sensors regularly for signs of wear, damage, or contamination that may affect performance. Recommendations regarding storage conditions, battery maintenance, and environmental protection are also included.

The troubleshooting section assists users in identifying and resolving common operational issues. Guidance is provided for situations involving low battery indications, sensor disconnection, signal loss, poor signal quality, unexpected measurement values, and alarm conditions. Corrective actions are described to facilitate restoration of normal device operation and reduce unnecessary service interventions.

The IFU also contains technical information relating to device specifications, measurement ranges, accuracy characteristics, operating conditions, and environmental limitations. This information enables users to understand the performance capabilities of the device and ensures that measurements are obtained under appropriate operating conditions.

Overall, the Instructions for Use provide comprehensive guidance that supports the safe and effective use of the pulse oximeter throughout its operational life. The information supplied within the IFU contributes to risk reduction, promotes measurement reliability, and assists users in obtaining clinically meaningful results while maintaining patient safety. The document forms an essential component of the overall risk management strategy and supports compliance with applicable regulatory requirements governing information supplied by the manufacturer.

Got it. For a Technical Documentation File, you wouldn't write about what a PMS Plan is. You would actually write the PMS Plan itself for the device. The same applies to the PMCF Plan.

Below is a CER/Technical File style version written directly for a Pulse Oximeter.

11. Post-Market Surveillance (PMS) Plan

11.1 PMS Plan for Pulse Oximeter

Device Information

The PMS activities described within this plan apply to the Pulse Oximeter intended for non-invasive measurement of arterial oxygen saturation (SpO₂) and pulse rate in adult, pediatric, and neonatal patients within the intended use environments specified by the manufacturer.

The device is intended to provide healthcare professionals and users with real-time physiological information to support patient monitoring and clinical decision-making. The PMS system has been established to continuously evaluate the safety, performance, usability, and benefit-risk profile of the device throughout its commercial lifecycle.

PMS Objectives

The objectives of the PMS program are:

- To confirm the continued safety and performance of the pulse oximeter during routine clinical use.

- To verify that the device continues to achieve its intended purpose.

- To identify previously unknown hazards and emerging risks.

- To monitor known side effects and residual risks identified in the Risk Management File.

- To evaluate trends associated with complaints, adverse events, and device malfunctions.

- To determine whether corrective or preventive actions are required.

- To verify that the benefit-risk profile remains acceptable throughout the device lifecycle.

- To provide ongoing input into Clinical Evaluation activities and Risk Management processes.

Data Sources

Post-market data shall be collected from multiple sources to ensure comprehensive monitoring of device performance.

The following information sources will be reviewed:

Complaint Handling System

All complaints relating to:

- Inaccurate SpO₂ measurements

- Incorrect pulse rate readings

- Sensor failures

- Battery failures

- Device shutdowns

- Display malfunctions

- Alarm failures

- Communication errors

- Mechanical damage

will be documented, investigated, and analyzed.

Service and Repair Records

Information collected from:

- Returned products

- Repair activities

- Technical support reports

- Preventive maintenance activities

will be evaluated for recurring technical issues.

Vigilance Reports

Monitoring shall include review of:

- Serious incidents

- Adverse events

- Regulatory reporting activities

- Field Safety Corrective Actions (FSCAs)

- Safety notices

Any reported incidents involving patient harm or potential patient harm will be investigated according to established vigilance procedures.

Literature Surveillance

Scientific literature searches will be conducted annually using:

- PubMed

- Embase

- Scopus

- Cochrane Library

Searches will focus on:

- Pulse oximeter safety

- Pulse oximeter performance

- Measurement accuracy

- Hypoxemia detection

- Skin pigmentation effects

- Low perfusion conditions

- Motion artifacts

- Neonatal applications

Relevant publications will be reviewed to identify new safety information or performance concerns.

User Feedback

Feedback received from:

- Healthcare professionals

- Biomedical engineers

- Hospitals

- Distributors

- Home healthcare providers

will be evaluated to identify usability issues and opportunities for product improvement.

Trend Analysis

Trend analysis shall be conducted annually.

The following indicators will be monitored:

- Complaint rates

- Sensor-related failures

- Battery-related failures

- Software-related events

- Alarm malfunctions

- Device return rates

- Measurement accuracy complaints

Any statistically significant increase in occurrence rates will trigger further investigation.

Risk Management Interface

PMS findings will be reviewed against the Risk Management File.

The manufacturer will evaluate whether:

- Existing risks remain acceptable.

- New hazards have emerged.

- Risk control measures remain effective.

- Additional mitigation measures are required.

Updates to the Risk Management File shall be implemented when necessary.

Clinical Evaluation Interface

PMS findings will be reviewed as part of the Clinical Evaluation process.

The manufacturer will assess whether:

- Clinical benefits remain unchanged.

- New safety concerns have emerged.

- Additional clinical evidence is required.

- Benefit-risk conclusions remain valid.

Where necessary, the Clinical Evaluation Report will be updated.

PMS Deliverables

Outputs generated from PMS activities include:

- Complaint Trend Reports

- Vigilance Assessments

- CAPA Records

- Literature Surveillance Reports

- PMS Reports

- Risk Management Updates

- Clinical Evaluation Updates

PMS Conclusion

Based on the established PMS process, continuous monitoring of real-world device performance will be maintained throughout the commercial lifecycle of the pulse oximeter. The collected information will be used to verify continued safety, performance, and clinical effectiveness while ensuring that the benefit-risk profile remains favorable for the intended patient population.

12. Post-Market Clinical Follow-Up (PMCF) Plan

12.1 PMCF Plan for Pulse Oximeter

Device Information

This PMCF Plan applies to the Pulse Oximeter intended for non-invasive measurement of oxygen saturation (SpO₂) and pulse rate.

The device utilizes optical sensing technology and is intended for use in healthcare facilities and other environments specified within the Instructions for Use.

Although pulse oximetry represents a well-established technology with substantial clinical evidence, PMCF activities will be conducted to continuously confirm the clinical safety, performance, and benefit-risk profile of the device under real-world use conditions.

PMCF Objectives

The objectives of PMCF activities are:

- To confirm the clinical performance of the pulse oximeter during routine use.

- To verify continued measurement accuracy.

- To confirm that the device continues to achieve its intended clinical purpose.

- To identify previously unknown risks.

- To assess whether residual risks remain acceptable.

- To monitor long-term safety and performance trends.

- To confirm the continued acceptability of the benefit-risk profile.

- To identify opportunities for product improvement.

PMCF Questions

The PMCF activities are designed to answer the following questions:

- Does the pulse oximeter continue to provide clinically acceptable oxygen saturation measurements during routine use?

- Does the pulse oximeter continue to provide clinically acceptable pulse rate measurements?

- Have any new safety concerns emerged since commercialization?

- Are the identified residual risks still acceptable?

- Does the benefit-risk profile remain favorable?

- Are there any new patient populations or use conditions requiring additional investigation?

PMCF Methodology

PMCF data will be collected using several complementary methods.

Scientific Literature Review

Annual literature reviews will be conducted using:

- PubMed

- Embase

- Cochrane Library

- Scopus

The review will focus on:

- Clinical performance

- Measurement accuracy

- Safety outcomes

- Adverse events

- Skin pigmentation effects

- Low perfusion studies

- Motion artifact studies

- Neonatal monitoring

- Home monitoring

- Remote patient monitoring

All relevant publications will be critically appraised and incorporated into PMCF evaluations.

Post-Market Clinical Data Review

Clinical information collected through PMS activities will be reviewed for:

- Device-related complaints

- Measurement inaccuracies

- Adverse events

- Serious incidents

- User-reported clinical concerns

The objective is to identify trends potentially impacting safety or performance.

User Experience Evaluation

Clinical feedback from healthcare professionals will be collected and reviewed.

Areas of interest include:

- Ease of use

- Reliability

- Alarm performance

- Sensor usability

- Clinical usefulness

- Interpretation of results

Findings will be incorporated into PMCF assessments.

State-of-the-Art Monitoring

PMCF activities will include ongoing monitoring of:

- Clinical guidelines

- International recommendations

- New pulse oximetry technologies

- Emerging performance standards

This information will be used to confirm that the device remains aligned with current state-of-the-art expectations.

PMCF Evaluation Criteria

The pulse oximeter will be considered clinically acceptable if:

- No unacceptable safety risks are identified.

- No significant increase in adverse events occurs.

- Clinical benefits remain unchanged.

- Performance remains consistent with validated specifications.

- Benefit-risk determination remains favorable.

PMCF Outputs

PMCF activities will generate:

- PMCF Evaluation Reports

- Literature Review Reports

- Clinical Performance Assessments

- Benefit-Risk Updates

- Clinical Evaluation Updates

- Recommendations for CAPA activities if required

PMCF Conclusion

The PMCF program has been established to provide continuous clinical evidence supporting the safety, performance, and effectiveness of the pulse oximeter throughout its post-market lifecycle. Through systematic literature review, analysis of post-market clinical information, user feedback evaluation, and state-of-the-art monitoring, the manufacturer will verify that the device continues to provide clinically meaningful benefits while maintaining an acceptable level of risk. The PMCF activities will support ongoing regulatory compliance and ensure that the Clinical Evaluation Report remains current and supported by up-to-date clinical evidence.

13. Declaration of Conformity- attached